PPGPT Anomaly Detection User Guide From Firebase vitals to AI-powered cardiac anomaly detection

This guide explains how to install, configure, and run the PPGPT anomaly detection pipeline. No prior AI or signal processing knowledge is required.

# 1.  What This Tool Does

This tool reads heart rate and blood oxygen data from a Firebase JSON file, builds a synthetic PPG waveform from those numbers, and feeds it into a pre-trained AI model (PPGPT) that was trained on thousands of real heartbeat signals. The model then flags whether the signal looks normal or anomalous.

The pipeline in plain language:

| Step | What happens | Output |
| --- | --- | --- |
| 1 | Read Firebase JSON | BPM + SpO2 values |
| 2 | Synthesise IR PPG wave | 500-sample waveform |
| 3 | Tokenise the wave | Integer tokens (0-101) |
| 4 | Run PPGPT model | Next-token predictions |
| 5 | Compute MAE score | Single error number |
| 6 | Apply threshold | NORMAL or ANOMALY |

# 2.  Project Structure

After setup, your folders should look like this:

Syncope_HeartRate/

HeartGPT/                  <- original library (do not edit)

Model_files/

PPGPT_500k_iters.pth  <- model weights (download separately)

core/  configs/  scripts/  ...

project/                   <- your working folder

main.py                <- the detection script

vitals.json            <- Firebase sensor data

ppgpt_anomaly_result.png  <- diagnostic plot (generated)

# 3.  Requirements

## 3.1  Python packages

Install all dependencies with:

pip install torch numpy matplotlib scipy

## 3.2  Model weights

The PPGPT weights file is not included in the repository because it is too large for GitHub (>100 MB). Download it from the HeartGPT repository and place it at:

HeartGPT/Model_files/PPGPT_500k_iters.pth

Download link

https://github.com/harryjdavies/HeartGPT

Go to: Model_files/ -> PPGPT_500k_iters.zip -> extract the .pth file

# 4.  Input Format

The script reads a JSON file produced by Firebase. The expected structure is:

{

"users": [

null,

{

"vitals": {

"BloodOxygen": 94,

"HeartRate":   89

}

}

]

}

The script also accepts a simpler flat structure:

{ "HeartRate": 89, "BloodOxygen": 94 }

# 5.  How to Run

## 5.1  Open a terminal in the project folder

cd C:\Users\yourname\Documents\Syncope_HeartRate\project

## 5.2  Basic run

Uses vitals.json in the current folder with default settings:

python main.py

## 5.3  Specify a custom JSON file

python main.py --input path/to/your/file.json

## 5.4  Save a diagnostic plot

python main.py --input vitals.json --plot

This saves a PNG file showing the synthesised waveform, token comparison, and per-token error chart.

## 5.5  Change the anomaly threshold

python main.py --input vitals.json --threshold 13.0

## 5.6  All options at once

python main.py --input vitals.json --threshold 13.0 --plot --plot-out result.png

# 6.  Command-Line Arguments

| Argument | Default | Description |
| --- | --- | --- |
| --input / -i | vitals.json | Path to Firebase JSON file |
| --model / -m | ../HeartGPT/Model_files/...pth | Path to PPGPT weights file |
| --threshold / -t | 6.5 | MAE threshold (tokens). Scores above = anomaly |
| --fs | 100.0 | Sampling frequency for waveform synthesis (Hz) |
| --plot | off | Save a diagnostic PNG when present |
| --plot-out | ppgpt_anomaly_result.png | Filename for the diagnostic plot |

# 7.  Understanding the Output

## 7.1  Terminal output

A typical run prints:

[vitals]      Heart Rate  : 89.0 bpm

[vitals]      Blood Oxygen: 94.0%

[clinical flags]

* Normal heart rate: 89 bpm

* Mild hypoxaemia  - SpO2 94%  (normal: >=95%)

[inference]   MAE score : 10.6072 tokens

Threshold : 6.5 tokens

+---------------------------------------+

|  ANOMALY DETECTED  (MAE = 10.61)      |

+---------------------------------------+

## 7.2  Clinical flags

These are rule-based checks applied directly to the raw vitals, independent of the AI model:

| Condition | Trigger | Meaning |
| --- | --- | --- |
| Bradycardia | BPM < 60 | Heart rate too low |
| Tachycardia | BPM > 100 | Heart rate too high |
| Mild hypoxaemia | SpO2 < 95% | Borderline low oxygen |
| Severe hypoxaemia | SpO2 < 90% | Critically low oxygen |
| Normal | All in range | No clinical concern |

## 7.3  The diagnostic plot (--plot)

Three panels are shown:

- Synthesised IR PPG wave — the heartbeat shape built from BPM and SpO2

- Token comparison — blue line (input) vs orange dashes (model predictions). Close tracking = normal signal

- Per-token error — spikes above the MAE line indicate positions where the model was most surprised

# 8.  The Anomaly Threshold

Important — calibrate before trusting results

The default threshold of 6.5 is a starting point, not a clinically validated value.

Because the input is a synthesised waveform (not a real sensor recording), the MAE

will always be higher than it would be with real PPG data. You must calibrate.

## How to calibrate

- Collect 5-10 readings from a healthy person (BPM 60-100, SpO2 >= 95%)

- Run each through the script and note the MAE scores

- Calculate: threshold = average MAE + 2x standard deviation

- Update the default in parse_args or pass it with --threshold

Example: if healthy readings give MAE scores of 10.2, 11.1, 10.8, 10.5 — average is 10.65, std is ~0.35, so a good threshold is around 11.4.

# 9.  Changing the Model Path

If you move files around, update the default path inside parse_args in main.py:

p.add_argument(

"--model", "-m",

default="../HeartGPT/Model_files/PPGPT_500k_iters.pth",

...

)

Or pass it directly at runtime without touching the code:

python main.py --model C:/full/path/to/PPGPT_500k_iters.pth

# 10.  Troubleshooting

| Error message | Cause | Fix |
| --- | --- | --- |
| JSON file not found | Wrong --input path | Check the path to your JSON file |
| Model weights not found | Missing .pth file | Download from HeartGPT repo and place in Model_files/ |
| Could not find HeartRate key | Wrong JSON structure | Check JSON matches the format in Section 4 |
| ModuleNotFoundError: torch | PyTorch not installed | Run: pip install torch |
| MAE always above threshold | Threshold not calibrated | Follow calibration steps in Section 8 |

# 11.  For Colleagues Cloning the Repository

- Clone the repo from GitHub

git clone https://github.com/YOUR_USERNAME/YOUR_REPO.git

- Download the model weights from the HeartGPT repo and place at:

HeartGPT/Model_files/PPGPT_500k_iters.pth

- Install Python dependencies:

pip install torch numpy matplotlib scipy

- Navigate to the project folder and run:

cd project

python main.py --input vitals.json --plot

PPGPT Anomaly Detection Pipeline  —  Syncope HeartRate Project